Physics 4B

Lab Quiz 2: The Oscilloscope

Overview of a oscilloscope:

An oscilloscope is an electronic measuring device which provides a two-dimensional visual representation of a signal. Because the oscilloscope allows the user to see the signal(s), their characteristics can be easily measured and observed. The oscilloscope (usually) displays a graph of voltage (on the vertical axis) versus time (on the horizontal axis). Most electrical circuits can be easily connected to the oscilloscope typically with probes.

· There are four major control groups on the Tektronix oscilloscope: (1) the Display group; (2)the Vertical group; (3) the Horizontal group; and (4) the Trigger group. See Figure 1, below. 
[image: image1.wmf]
Figure 1: Front Panel of Analog Oscilloscope

1. Display Group

This group is used to display and adjust the signal for optimal viewing. It consists of the display screen, the intensity-control knob, the beam-find button, the focus-control knob, trace rotation screw, and the power switch. 


The display screen is laid out in a 8 by 10 centimeter grid. The oscilloscope draws a ‘trace’ or graph by moving an electron beam across the phosphor coating on the inside of the cathode-ray tube (CRT).  The excited phosphorous glows for a short period of time, thereby tracing the path of the beam.


The intensity-control knob is used to adjust the brightness of the trace. The level should be set to the user's preference; however, increasing the intensity beyond a certain point will make the trace "fuzzy." The intensity level should never be increased past the point where the trace has sharply defined edges.


The beam-find button allows the user to locate the electron beam anytime it's off-screen. Push the beam find button to temporarily reduce the vertical and horizontal deflection voltages so that the beam always appears within the 8 by 10 centimeter screen.  

· The trace rotation screw permits alignment of the trace with the horizontal graticule lines.


The focus-control knob adjusts the electron beam for optimal trace resolution. 
2. Vertical Group:

This group is used to adjust the vertical components (Y-axis) and the vertical position of the signal. This group consists of the vertical-position knobs, channel-selector switch, volts-per-division-selector knobs, variable volts/div (CAL) knob, input-coupling switch, the channel-mode-selector switch, the channel-2-invert switch, and the BNC connectors. Because the Tektronix 2201 oscilloscope is a two channel oscilloscope, there is one set of switches for each channel.


The vertical-position controls are used to vertically move the trace of one channel or the other.  


The channel-selector switch, labeled ‘CH1/BOTH/CH2’, selects which channels are displayed on the screen.


The volts-per-division-selector knob sets the vertical scale for each channel's trace. 

a. Each channel can have a different vertical scale.

b. A division is one "block" on the screen.

c. Each tick mark is one-fifth (0.2) of a division.

d. This knob is marked in both volts and millivolts.

· The variable volts/div (CAL) knob provide uncalibrated variable deflection factors between the calibrated settings of the volts/div controls. The volts/div sensitivity is reduced by at least 2.5 times at the fully counterclockwise rotation of the variable knob. The detent at full clockwise rotation is the CAL position of the variable knob.


The input-coupling switch, labeled ‘AC/GND/DC’, selects the coupling mode of that channel's display.  AC means that only the alternating portion of the signal is displayed.  DC will display both the alternating portion of the signal, plus any DC component.  GND shows the 0 V reference level.


The channel-mode-selector switch, labeled ‘ADD/ALT/CHOP’, is activated only when BOTH is selected on the channel-selector switch.  ADD graphically adds the Channel 1 signal to the Channel 2 signal.  If the channel-2-invert switch (‘NORMAL/INVERT’) is set to INVERT, then ADD will in fact subtract the Channel 2 signal from the Channel 1 sigal.  ALT traces one channel, then the next.  CHOP works like ALT, but jumps back and forth between the two channels during a single trace.

3. Horizontal Control Group:

The Horizontal Control Group consists of the coarse and fine position knobs, the horizontal-magnification switch, the seconds-per-division-selector knob, and the variable sec/div knob (CAL). 


The position-knobs allow the horizontal movement of the traces. 
· The horizontal-magnification switch (‘X1/X10’) selects the regular (X1) and a horizontally magnified (X10) trace.


The seconds-per-division knob sets the time base (horizontal) scale.  It is marked in seconds, milliseconds, and microseconds.

Note: There is only one horizontal scale for both channels.

· The variable sec/div knob (CAL) continuously varies the sweep time per division to at least 2.5 times the calibrated time per division set by the sec/div switch. Full counterclockwise rotation increases the slowest sweep time per division to at least 2s.

4.  Trigger Group

In order to display a signal, the oscilloscope must be able to ‘lock’ onto that signal; the function of the triggering controls is to do just that.  Triggering can be a complicated topic, and is beyond the scope of this lab.  For more details, the student is referred to the Tektronix manual.  The trigger group is made up of the trigger level knob, the rising/falling edge switch, the trigger mode switch, and the trigger source switches.

The trigger-level knob sets the voltage level at which the oscilloscope will ‘trigger’.  If a signal is ‘running’ -- that is, not stable -- the trigger level may be too high or too low for the oscilloscope to recognize the signal.  Often you can lock in a running signal by rotating this knob to the left and to the right. This adjustment provides a mechanism for ignoring small (low voltage) signals that are well below the level of the signal you ar interested in.  This knob is not indexed.


The rising/falling-edge switch selects whether the oscilloscope will trigger on the positive (rising) or negative (falling) edge of the signal.


The trigger-mode switch will normally be set to AUTO, but sometimes it is necessary to use NORM.  Other settings are not relevant to this lab.


The trigger-source switches select which signal the oscilloscope will attempt to lock onto.  Possible choices include CH1, CH2, VERT MODE, or EXTERNAL.   Selecting CH1 or CH2 will make the oscilloscope attempt to trigger on those channels.  If no input is available on that channel, there may be problems when attempting to view both channels simultaneously.  For this reason, it is recommended that this switch be left on VERT MODE, which provides an automatic trigger on either CH1 or CH2.

How To Make Readings on the Oscilloscope:


Peak-to-Peak Voltage:

a. Use the vertical-position knob to place a peak (positive or negative) on a horizontal line, keeping the peak on the screen.

b. Use the horizontal-position knob to set the next (opposite sign) peak on the center vertical line.

c. Count the number of divisions between the positive and negative peaks.

Note: The divisions are further subdivided into fifths (0.2) by the tick marks.

d. Multiply the number of divisions from step 3 by the volts/div setting for the channel in use.


Period:

a. Use the horizontal-position knobs to align any edge of the signal with a vertical line.

b. Use the vertical-position knob to place the next identical edge crossing on the X-axis.

c. Count the number of divisions along the horizontal line to the next crossing in the same direction.

Note:  The divisions are subdivided into fifths (0.2) by the ‘tick marks’.

d. Multiply the number of divisions from step 3 by the sec/div setting.


Frequency:

Measure the period of the signal (T).

Calculate frequency (f) using  f = 1/T.

[image: image2.png]TEKTRONIX 2201

VERTICAL HORIZONTAL TRIGGER
INTENSITY | posirion PoSITION | POSITION LEvEL
BEAM FIND NORM INVERT

CHIBOTH CH2, ADD ALTCHOP|  MAG MODE

FOUS | Gr voLTSIDIV  CH2VOLTS/DIV | SECIDIV | SOURCE

POWER ACoNDDC ACGNDDC
cHi e B

O I'e) o O





Figure 2: Measuring a Signal.

· Example:

a. Assuming the Volts/Div knob reads 2 V/div, the above peak-to-peak voltage would be: 

Vpp = 2 volts/div x 5.2 div = 10.4 volts

b. Assuming the Sec/Div control knob read 50ms, the above period would be: 

T = 50 milliseconds/div x 5.25 div =  2625.0 milliseconds = 2.625 seconds

Practice Procedure:

For practice and during the quiz, you will use the function generator to generate a signal and then use the analog oscilloscope to make some measurements of the signal.  The quiz will consist of doing the same measurements that are asked for at the end of this hand-out.  However, you will not be able to refer to these notes and Ms Shang or Dr. Harris will choose the settings on the power supply and will “sabotage” the oscilloscope by randomly adjusting various knobs and settings.  Your challenge will be to re-obtain the oscilloscope display and do the necessary measurements.   The following checklist will not be available to you during the quiz.
· Connect the function generator output to the oscilloscope Channel 1 input with a BNC cable, and make sure all the knobs on the function generator are pushed in. 

· Select a triangle wave on the generator with the frequency around 1.0 kHz, and set the amplitude knob to 12 o’clock.  (For additional practice, try other settings on the power supply.
· Turn on the power on both instrument.
· Set the oscilloscope CH1/BOTH/CH2 switch to CH1, AC/GND/DC switch to GND. Use the CH1 vertical position knob to move the ground (0V) reference (horizontal line) to the center line on the screen.
· Adjust the signal intensity and focus to a comfortable level, by using the intensity and focus knobs, respectively.
Note: Keep the signal intensity to within a reasonable range to minimize the chances of burning the phosphor and damaging the display screen.
· After referencing the signal to ground, set the AC/GND/DC switch to AC, and set the Volts-per-Division knob to 2 V/div.
· Set the NORMAL/INVERT switch  to NORMAL, and ADD/ALT/CHOP switch to ALT.
· Set the MAG switch to X1, and the seconds-per-division knob to 0.2 ms/div.
· Set the rising/falling-edge switch to positive, the trigger-mode switch to AUTO, and the trigger-source switch to VERT MODE.

· Make sure that the 'cal' knobs on the V/div and S/div knobs are both turned to the fully clockwise (CW) position.  They will click, or lock, into position ensuring that the scope will be calibrated properly.
Physics 4B
Oscilloscope Quiz

Name:_________________________




Score: _________

The examiner will set the power supply functions.  Your assignment is to obtain a stable display of the signal and perform the following measurements.  

I.  Describe the first signal:_________________________________ 
a. What is the setting on the Volts/Div control knob?  _______ volts/div

b. How many vertical divisions from peak-to-peak?  _______ div

c. What is the peak to peak voltage (Vpp)?



Vpp =  _______ volts/div  x  _______ div =  _______ volts

d. What is the setting on the Sec/Div control knob?  __________ seconds/div

e. How many horizontal divisions from positive going crossing to positive going crossing?  _______ div

f. What is the period of the signal (T)?



T =  _______ seconds/div  x  _______ div =  __________ seconds

g. What is the frequency of the signal (f)? __________ hertz

II.  Measure the frequency of the tuning fork provided using a microphone.

a) What is the setting on the Sec/Div control knob?  __________ seconds/div

b) How many horizontal divisions from positive going crossing to positive going crossing?  _______ div

c) What is the period of the signal (T)?


T =  _______ seconds/div  x  _______ div =  __________ seconds

d) What is the frequency of the signal (f)? __________ hertz

Examiner:_____________________________

Date:______________________

































































































6

